Models for synchrophasor with step discontinuities in magnitude and phase: estimation and performance

Marcelo B. Martins, Renata de Barros e Vasconcellos, and Paulo A. A. Esquef, Member, IEEE

Abstract  —  This work proposes an alternative method to assess the calibration of phasor measurements units (PMUs) under  conditions of step discontinuities in magnitude or phase. Two parametric mathematical models, which are proposed to  represent these signals, are fitted to signal samples via an iterative numerical method. The proposed approach does not require any time adjustment of the analysis window to skip the discontinuity. The estimated parameters can be used to calculate a reference phasor with an appropriate definition. Moreover, accurate estimates of the location and magnitude of the discontinuities are provided.

Index Terms — Calibration, dynamic tests, phasor measurement units, synchrophasor, uncertainty.

# INTRODUCTION

T

he dynamic behavior of modern electric grid demands testing the performance of phasor measurement units (PMUs) under magnitude steps and phase steps [1]. The accuracy of those measurements depends on the reference values provided by PMU calibration systems. Recent developments towards the calibration of PMUs for distribution grids demand lower uncertainty levels than the first systems, which were designed for the context of transmission grids [2]. The calibration process depends on generating and sampling synchronized waveforms, from which reference phasors are compared to the values provided by the PMU under test. The phasor which phase is related to an ideal cosine function with maximum centered in the Universal Time Code (UTC) second is called a synchrophasor. PMUs also provide estimates of frequency and rate of change of frequency (ROCOF) at a given report rate.

For the estimation of synchrophasors, a stationary phasor waveform can be curve fitted with a steady state sinusoidal function with good accuracy. Methods to estimate parameters of signals with slowly varying frequency, magnitude or phase, for PMU calibration purposes, are also presented in [2] and [3]. Typical variations in parameters or nonlinearities can be modeled by low order Taylor series expansion. Then, iterative procedures are used to estimate the model parameters.

In the specific case of an observed phasor disturbed by a step discontinuity in magnitude or phase, the estimation using an underlying steady state model is inappropriate and does not guarantee convergence nor accuracy. Besides, there is a lack of definition of what the reference phasor should be. To overcome this difficulty, the method used in [3] adjusts the timestamp and position of the analysis window to skip the discontinuity and set the phasor estimates where the discontinuity occurs with those of obtained from the previous or following window. That way, it avoids the mathematical modelling of a step discontinuity and considers the reference value coming from the steady state.

The mentioned procedures are designed to calibrate PMUs, and may be not detailed enough to evaluate the performance of calibration systems. Methods for a more detailed analysis of calibration systems under step conditions are proposed in [4]. The authors use a pointwise root mean squared error for the performance evaluation of the investigated phasor estimators.

Although not trying to break the stationary paradigm of the phasor representation of a signal, we propose an alternative approach to evaluate the measurements of PMU calibration systems. For that, this extended version of [5] offers the following contributions:

1) Aiming at having accurate phasor estimates under transient conditions, we propose signal models that account for step discontinuities in magnitude or phase, incorporating representations of both the transient and the phasor.

2) Ways to estimate the parameters:

a) To use the instantaneous frequency as provided by the Hilbert transform of the perturbed phasor signal to estimate the instant of the step discontinuity;

b) To use a nonlinear least-square method (NL-LS) to estimate the other parameters of the models proposed in 1), provided the step instant estimate.

3) Proposition of single phasor estimate for transient situations, which can be used as reference values for PMU calibrations and easily implemented in the existing systems, in place of traditional estimation schemes. The proposed method depends on the estimation of parameters of an underlying model.

4) A preliminary evaluation of a laboratory system intended to be a PMU calibrator based on the aforementioned methods.

Fig. 1  Transitioning phasors for a) magnitude step, b) phase step

In order to assess the contribution to the uncertainty of the calibration system, we made simulations to obtain the numerical errors of each numerical method. In addition to this introduction, the paper is organized as follows: in section II, we present the mathematical background related to the proposed models for dynamic signals with magnitude or phase steps; the intermediary phasor definitions; and the basic concepts of the Hilbert transform and of the NL-LS method. In section III, we describe the Monte Carlo simulations run to analyze the numerical errors of each method. In section IV, we detail the Laboratory measurements devised to evaluate the use of the proposed system for PMU calibration. Finally, in section V, we  discuss the attained results and draw conclusions on the reported investigation.

# Mathematical Background

## Mathematical models for dynamic signals

A pure sinusoidal waveform with one magnitude step, located at t=τ can be modeled in continuous time

yt=x11+x2ut-τcosωt+φ+ ηt,    (1)

where ut is the step function. A similar model for the phasor waveform with one phase step is

yt=x1cosωt+φ+ x3ut-τ+ ηt,        (2)

where the step function ut-τ is used as an idealization of a fast transient in magnitude or phase occurring at the instant τ, where x1 is the signal nominal magnitude, x2 is a decimal value representing the magnitude change, x3 is the amplitude of the phase step, ω is the angular frequency, φ is the initial phase, and ηt represents interfering noise. Provided a sufficiently accurate estimate of τ, the set of parameters P=x1,x2,x3, ω,φ can then be adjusted to obtain a waveform that best fits the data received by the calibration system sampler,. Given a prescribed signal to noise ratio (SNR) in dB, for a zero mean gaussian white noise, the variance of noise is

η0=σy10SNR202,                                (3)

where σy is the standard deviation of the signal yt.

## Reference phasor values

After one estimates the model parameters, the problem of obtaining one phasor that represents the waveform arises. Instead of considering the values estimated from the analysis windows adjacent to the transient, one alternative proposal could be an intermediate value for magnitude or phase. The concept is illustrated in Fig. 1, where the phasor V1 represents the waveform during an initial steady state, Ve is a phasor that could be possibly representative of an intermediate state during the occurrence of a magnitude or phase step, and V2 represents the signal in the final steady state condition. (In Fig. 1-a), Ve is taken off the axis only for visualization purposes.)

Intermediate phasor estimates can be obtained, for example, using a weighted means out of the estimated model parameters. Using the notation x to represent an estimate of x, for any τ∈0,T, for a waveform with magnitude step described by (1), the estimated intermediate phasor would be

Ve= Xe∠φ=x1τ+x11+x2T-τT∠φ ;          (5)

and for a waveform with a phase step test described by (2), the intermediate phasor would be

Ve=X∠φe=X∠φτ+φ+x3T-τT   .               (6)

## Instantaneous frequency via Hilbert transform

Hilbert transform has been used to estimate instantaneous frequency (IF) of narrowband monocomponent signals, which is the case of ideal electric network phasor components. There are various applications of IF estimation reported in the literature, e.g., characterization of electric disturbances [6] and detection of edits in audio signals that bear the electric network frequency [7]. Anomalous perturbations on the IF can flag the occurrence of discontinuities in the signal. The time instant they happened can be estimated via appropriate amplitude threshold schemes. The concept is shown in Fig. 2 and Fig. 3.

Given a real narrowband monocomponent signal xt, -∞<t<∞, let zt be called the analytic signal associated to xt, defined as

zt=xt+jxt,                              (7)

where

xt=Hxt=-∞∞xuπt-udt,              (8)

is the Hilbert transform of xt. If zt is expressed in the polar form

zt=Atejθt,                              (9)

At= x2t+x2t,                     (10)

θt= tan-1xtxt,               (11)

the instantaneous frequency (IF) can be defined as

fi= 12πdθtdt.                            (12)

The discrete implementation of H[xt)] can be done with the fast fourier transform (FFT). [ref]

Fig. 2 - Instantaneous frequency of signal with one magnitude step

Fig. 3 - Instantaneous frequency of signal with one phase step

## Model Estimation via Levenberg-Marquardt

Fig. 4 - Histogram of errors in step instant estimation

Consider N samples from a sequence yk, which can be either generated via computational simulation or sampled from measured real phenomenon, with uniform sampling period Δt. One wishes to fit the models (see Section II) with parameters P to yk. For that, one can define the error cost function

εP=12k=1Nyk-yk∆t2,                 (4)

and try to solve the minimization problem minPεP.

The estimation of phasor parameters considering variations in frequency within the model requires dealing with a non-linear function. Existing calibration systems solve this problem for steady-state signals through low order Taylor linearization [3], or using directly some non-linear minimization algorithm, e.g., Levenberg-Marquardt (LM) [4].

The Levenberg-Marquardt (LM) algorithm is an iterative technique for nonlinear minimization problems. It combines the Gauss-Newton method and the steepest descent, being very useful when the size of the algorithm step cannot be obtained in a closed form. Such NL-LS methods can reach local minima and need a convex cost function.

Explicar mais um pouco e colocar referencia

Equações

Criterios de convergencia

Aproximação numérica da derivada por diferenças centrais

# Numerical Simulations

We performed Monte Carlo simulations with 1000 runs for each set point, to estimate the errors obtained with the numerical computation of the reference values. These errors will be considered as an uncertainty component in the overall process.

The input signals were digitally generated, with all nominal parameters prescribed in the standard [1] and random values representative of expected uncertainties in each parameter. The signals were created based on (1) and (2) with a 5 kHz sampling frequency, with a duration of 0.1 s. The nominal values are summarized in Table I, with their respective uncertainties used in simulations.

TABLE I

Nominal values and uncertainties for simulations

| Parameter | x1 | x2 | x3 | ω2π | φ |
| --- | --- | --- | --- | --- | --- |
| Nominal | 1 Vpp | ± 0.1 | ± 10° | 60 Hz | 360°, ±120° |
| U[%] | 1 | 1 | 1 | 0.05 | 1 |

## Step instant estimation with Hilbert transform

The peaks were detected taking the maximum value of the IF subtracted by its median value. If we use an ideal signal without uncertainties in the parameters and nominal frequency, for a total duration of the window T, τ∈0.05T,0.95T, and SNR>40dB, the maximum absolute errors are not greater than 1∆t.

In a second simulation, designed to represent a more realistic situation, we allowed 500 ppm variation in frequency and 1% variation in the other parameters, all under a uniform distribution. The maximum errors obtained are not greater than 2∆t for a SNR>75dB. The distribution of τ errors (in Δt

Fig. 5 - Phase Step test: maximum errors for intermediate phasor

units) for positive magnitude step of 10% is shown in a histogram in Fig. 4. Similar histograms were obtained for negative magnitude steps and phase steps of ± 10°, with the same results.

## Parameters estimation with non-linear least squares

For the uncertainty analysis reported in this subsection, at each Monte Carlo run, the signal yk is generated with uncertainties added to the parameters, drawn from a uniform distribution centered in the nominal values, as shown in Table I.

In the iterative LM algorithm, the model parameters are initiated at the nominal values, and the optimization procedure seeks for the minimum point of εP,  which is reached at the actual values of the parameters.

Criterios de convergencia

The final estimates have significantly lower errors than the initial values, despite having some sensitivity to noise. As an example, the maximum errors obtained are shown in Table II for magnitude step tests and Table III for phase step tests. One can see that, for both tests, the estimation of the step discontinuity parameters are the most sensitive to the SNR. Moreover, they exhibit the highest maximum errors.

TABLE II

Maximum errors for magnitude step tests [ppm]

| SNR | x1 | x2 | ω2π | φ |
| --- | --- | --- | --- | --- |
| 97 | 1.1 | 19 | 0.07 | 0.34 |
| 95 | 1.6 | 22 | 0.11 | 0.49 |
| 93.5 | 1.6 | 27 | 0.11 | 0.49 |
| 92.5 | 1.7 | 31 | 0.13 | 0.58 |
| 90.5 | 2.7 | 42 | 0.19 | 0.83 |

TABLE III

Maximum errors for phase step tests [ppm]

| SNR | x1 | x3 | ω2π | φ |
| --- | --- | --- | --- | --- |
| 97 | 0.80 | 20 | 0.15 | 0.72 |
| 95 | 0.96 | 24 | 0.21 | 1.1 |
| 93.5 | 1.1 | 28 | 0.22 | 1.2 |
| 92.5 | 1.4 | 34 | 0.25 | 1.5 |
| 90.5 | 1.7 | 40 | 0.31 | 2.1 |

## Estimation of intermediate phasors

For signals digitally generated with the uncertainties indicated in Table I, and also considering maximum errors of ±2∆t in the estimation of τ, the intermediate phasors calculated with (5) and (6) were compared to nominal values to obtain magnitude and phase maximum errors.

For the magnitude step tests, with SNR<90.5 dB, in every prescribed situation, the maximum absolute errors obtained for the estimates of magnitude were not greater than 2.5 ppm, and for the estimates of phase the maximum absolute errors were not greater than 3.0x10-4 degrees. There is some sensitivity to noise, shown in Fig. 6, in function of the step instant of occurrence.

# Laboratory measurements

Aiming at validating the proposed method with real signals, several measurements were made using one digital sampling voltmeter (DSVM) and one arbitrary waveform generator (AWG), controlled by a personal computer (PC) via GPIB. The connections are shown in the block diagram of Fig. 7.

Fig. 6 – Magnitude Step test: maximum errors for intermediate phasor

The same waveforms used in simulations are reproduced by the AWG, with a nominal output of 1 Vpp, and sampled by the DSVM. Both are triggered with a 1 PPS (pulse per second) signal, coming from an atomic clock, so we can control the initial phase. The internal clock from the DVM is used as an external 10MHz reference signal by the generator. 5000 samples are taken during 1 s and stored in the DVM´s internal memory.

Fig. 6 - Block Diagram

Fig. 8 - Magnitude step occurrence in the 5th window

Fig. 9  Intermediary magnitude for 10% negative step.

The standard [1] establishes that the synchrophasors must be obtained related to the center of a window. Setting 500 samples/window, the first complete window will happen after 250 samples, after which we have 9 windows containing 6 cycles of 60Hz. The steps of magnitude or phase occur in the 5th window, as shown in Fig. 8. According to the procedure for equivalent sampling, the instants of occurrence of the steps are a set of equally spaced intervals τk=0.1kT, k=1…9.

For the windows with steady state waveforms, the same fitting algorithm used in [3] is used to obtain the synchrophasors and frequency estimates.

For the 5th window, the intermediate phasors were calculated using (5) or (6), after obtaining the parameters using the Hilbert algorithm for the step instant, and the-LM algorithm for the others. The frequency was obtained directly from the LM estimation.

The estimates of step instant were not greater than 2∆t, inside the expected uncertainty. The other parameters require a more detailed analysis.

## Intermediary magnitude and phase

The system is capable of providing reasonable intermediary values for magnitude and phase, as can be seen in Fig. 9 and

Fig. 10, respectively. Each series of data show different step instant occurrences in the 5th timestamp, in percentage of the window period T.

Fig. 10  Intermediary Phase. Phase positive step of 10 degrees, starting at 120 degrees.

Frequency

Frequency estimates for the windows with steady state signals have absolute variations not greater than 1.5 mHz from the nominal. When submitted to magnitude steps, the system has maximum frequency variations of about 9 mHz around the nominal value, as can be seen in Fig. 11, where the vales for each plot are taken with different step instant occurrences.

Fig. 11  Frequency variation during magnitude tests in function of the step instant.

The worst frequency variation related to the nominal happens when the system is submitted to a phase step, as shown in Fig. 12, where differences of about -23 mHz can be seen.

Fig. 12  Frequency variation during phase steps in function of the step instant.

## Magnitude during phase step

Fig. 13  Magnitude variation during phase steps in function of the step instant of occurrence

During steady state conditions, the magnitude have a variation around the average values of about 200 ppm. When submitted to phase steps, the magnitude differences to the average values obtained from the steady state phasors show higher values, as shown in Fig. 13, of about 300 ppm, with some maximum values reaching 400 to 500 ppm.

## Phase during magnitude step

Fig. 14  Phase variation during magnitude steps in function of the step instant of occurrence.

During steady state conditions, the phase estimates have variations around x degrees around the average values. The estimates of phase during magnitude step presents higher variations, especially when the step is not centered in half the window, even higher as far as the step occurrence is from the center of the window, as can be seen in Fig. 14. Phase errors can be of about 0.12 degrees, but it goes down to 0.02 degrees if τ=0.5T.

# Conclusion

Models for phasor signals disturbed by magnitude and phase step discontinuity were proposed, in the context of assessment of PMU calibration systems in transient conditions. Estimation of the model parameters via a nonlinear least-squares method was outlined. The proposed approach tackles the estimation of the step discontinuities in the phasor signal observed within an analysis window, instead of dodging the problem. Moreover, single phasor parameters are proposed for transient conditions.

The estimation accuracy of each parameter was obtained under different noise conditions and uncertainties forced upon the model used to generate the test signals. Cross-correlated uncertainties were found, which point out to the need of a deeper investigation and the possibility of further improvements in the estimation performance.

Within the limits reported, the proposed method can give reliable and accurate results to be used in PMU calibration systems, avoiding procedures to adjust windows and time-stamps around the instant of step occurrence.

Acknowledgment

I would like to thank Allen Goldstein, for the valuable comments and guidelines.

References

[1]	IEEE Standard for Synchrophasor Measurements for Power Systems, IEEE Std C37.118.1-2011, Dec. 2011

[2] D. Colangelo et al., "Architecture and characterization of a calibrator for PMUs operating in power distribution systems," 2015 IEEE Eindhoven PowerTech, Eindhoven, 2015.

[3]	Y. H. Tang, G. N. Stenbakken and A. Goldstein, "Calibration of Phasor Measurement Unit at NIST," in IEEE Transactions on Instrumentation and Measurement, vol. 62, no. 6, pp. 1417-1422, June 2013.

[4]	FRIGO, Guglielmo et al., “Definition of accurate reference synchrophasors for static and dynamic characterization of PMUs.” in IEEE Transactions on Instrumentation and Measurement, vol. 66, n. 9, p. 2233-2246, 2017.

[5]	ESTE MESMO

[6]	Caciotta, M., et al. "Detection of short transients and interruptions using the Hilbert transform." in Proc. 19th IMEKO World Congr. Fundam. Appl. Metrol.. 2009.

[7]	Edit Detection in Speech Recordings via Instantaneous Electric Network Frequency Variations

- Referencias de Guilherme e renata, na parte de incertezas estimadas

- Ver outras

First A. Author (M’76–SM’81–F’87) and all authors may include biographies. Biographies are often not included in conference-related papers. This author became a Member (M) of IEEE in 1976, a Senior Member (SM) in 1981, and a Fellow (F) in 1987.  The first paragraph may contain a place and/or date of birth (list place, then date). Next, the author’s educational background is listed. The degrees should be listed with type of degree in what field, which institution, city, state, and country, and year the degree was earned. The author’s major field of study should be lower-cased.

The second paragraph uses the pronoun of the person (he or she) and not the author’s last name. It lists military and work experience, including summer and fellowship jobs. Job titles are capitalized. The current job must have a location; previous positions may be listed without one. Information concerning previous publications may be included. Try not to list more than three books or published articles. The format for listing publishers of a book within the biography is: title of book (publisher name, year) similar to a reference. Current and previous research interests end the paragraph.

The third paragraph begins with the author’s title and last name (e.g., Dr. Smith, Prof. Jones, Mr. Kajor, Ms. Hunter). List any memberships in professional societies other than the IEEE. Finally, list any awards and work for IEEE committees and publications.  If a photograph is provided, it should be of good quality, and professional-looking. Following are two examples of an author’s biography.

Second B. Author was born in Greenwich Village, New York, NY, USA in 1977. He received the B.S. and M.S. degrees in aerospace engineering from the University of Virginia, Charlottesville, in 2001 and the Ph.D. degree in mechanical engineering from Drexel University, Philadelphia, PA, in 2008.

From 2001 to 2004, he was a Research Assistant with the Princeton Plasma Physics Laboratory. Since 2009, he has been an Assistant Professor with the Mechanical Engineering Department, Texas A&M University, College Station. He is the author of three books, more than 150 articles, and more than 70 inventions. His research interests include high-pressure and high-density nonthermal plasma discharge processes and applications, microscale plasma discharges, discharges in liquids, spectroscopic diagnostics, plasma propulsion, and innovation plasma applications. He is an Associate Editor of the journal Earth, Moon, Planets, and holds two patents.

Dr. Author was a recipient of the International Association of Geomagnetism and Aeronomy Young Scientist Award for Excellence in 2008, and the IEEE Electromagnetic Compatibility Society Best Symposium Paper Award in 2011.

Third C. Author, Jr. (M’87) received the B.S. degree in mechanical engineering from National Chung Cheng University, Chiayi, Taiwan, in 2004 and the M.S. degree in mechanical engineering from National Tsing Hua University, Hsinchu, Taiwan, in 2006. He is currently pursuing the Ph.D. degree in mechanical engineering at Texas A&M University, College Station, TX, USA.

From 2008 to 2009, he was a Research Assistant with the Institute of Physics, Academia Sinica, Tapei, Taiwan. His research interest includes the development of surface processing and biological/medical treatment techniques using nonthermal atmospheric pressure plasmas, fundamental study of plasma sources, and fabrication of micro- or nanostructured surfaces.

Mr. Author’s awards and honors include the Frew Fellowship (Australian Academy of Science), the I. I. Rabi Prize (APS), the European Frequency and Time Forum Award, the Carl Zeiss Research Award, the William F. Meggers Award and the Adolph Lomb Medal (OSA).

This paragraph of the first footnote will contain the date on which you submitted your paper for review. It will also contain support information, including sponsor and financial support acknowledgment. For example, “This work was supported in part by the U.S. Department of Commerce under Grant BS123456.”

The next few paragraphs should contain the authors’ current affiliations, including current address and e-mail. For example, F. A. Author is with the National Institute of Standards and Technology, Boulder, CO 80305 USA (e-mail: author@ boulder.nist.gov).

S. B. Author, Jr., was with Rice University, Houston, TX 77005 USA. He is now with the Department of Physics, Colorado State University, Fort Collins, CO 80523 USA (e-mail: author@lamar.colostate.edu).

T. C. Author is with the Electrical Engineering Department, University of Colorado, Boulder, CO 80309 USA, on leave from the National Research Institute for Metals, Tsukuba, Japan (e-mail: author@nrim.go.jp).